# RADFET ISS Dose-Analysis Methodology

Converting flight dosimeter telemetry to absorbed dose, with shielding comparison and uncertainty

Audience: R&D / Science team
Prepared: 2026-06-17
Status: Draft for review

This document describes how we will analyze the dosimetry data returned from the ISS. The flight data arrives in the same CSV format as LeadEnclosureData.csv. A runnable reference implementation of every step below lives in the sample_analysis/ folder and is demonstrated on synthetic data.

# 1. Purpose and scope

We fly five RADFET dosimeters, each behind a different shielding stack, to measure how effectively each stack attenuates the on-orbit radiation dose. This document defines the agreed procedure for turning the raw voltage telemetry into absorbed dose (in Rad), for combining repeat measurements, for quantifying uncertainty, and for comparing the shielding configurations.

# 2. Input data format

Each row is one channel read at one time. Columns (identical to LeadEnclosureData.csv):

| Column | Meaning |
| --- | --- |
| timestamp | ISO-8601 read time (host clock). |
| sensor_group | R1 or R2 - the two read-out groups (see Section 4). |
| channel | 1-5 - identifies the sensor / shielding configuration. |
| raw_adc | 12-bit ADC count (0-4095) behind the voltage. |
| voltage | Threshold-voltage shift dVt in Volts (already computed). |
| dose_rad | Convenience dose column; the analysis recomputes dose itself. |
| raw_timestamp | Secondary timestamp; may contain epoch corruption. |

Important: the voltage column already carries the threshold-voltage shift dVt (irradiated minus non-irradiated reference), so no extra baseline subtraction is performed in the dose conversion.

# 3. Calibration model

We use the Varadis QF 16 RADFET calibration (400nm IMPL RADFET, plastic package; Mask-set COMRAD; Lot P5925-W3; Co-60 source at 5.0 kRad/hour; Issue No.01, 01/08/2023). The fit relates the threshold-voltage shift to absorbed dose by a power law:

dVt [V] = A * Dose[Rad] ^ B

Inverting to recover dose from a measured dVt:

Dose[Rad] = ( dVt / A ) ^ (1 / B)

Varadis provides FIVE fits of the same data, each valid over a different dose range, and advises using the curve applicable to the actual dose. The narrow fits are the most accurate at low dose. Coefficients (single source of truth: calibration.py):

| Dose range | A | sigma(A) | B | sigma(B) | R^2 |
| --- | --- | --- | --- | --- | --- |
| 0-1 kRad | 0.0008 | 5.374e-05 | 0.9119 | 1.034e-02 | 1.000 |
| 0-5 kRad | 0.0033 | 2.994e-04 | 0.7062 | 1.104e-02 | 0.997 |
| 0-10 kRad | 0.0068 | 4.496e-04 | 0.6164 | 7.490e-03 | 0.996 |
| 0-50 kRad | 0.0241 | 5.361e-04 | 0.4752 | 2.150e-03 | 0.997 |
| 0-100 kRad | 0.0295 | 2.992e-04 | 0.4551 | 9.215e-04 | 0.999 |

# 4. Sensors, shielding, and trials

Five sensors are flown, one per channel, each behind a different shielding stack:

| Channel | Shielding configuration |
| --- | --- |
| 1 | None (bare) (reference for attenuation) |
| 2 | 2 mm Al |
| 3 | MLC1 |
| 4 | MLC1-b + Al |
| 5 | MLC2 |

Each sensor is read out at two positions/units recorded as sensor_group R1 and R2. We treat R1 and R2 as two independent TRIALS (replicates) of the same shielding configuration. This gives, per configuration, two dose estimates whose agreement is itself a check on repeatability.

# 5. Calibration-curve (range) selection

Because the correct curve depends on the dose we are trying to measure, we select it self-consistently. The five ranges are nested (0-1 kRad within 0-5 kRad within ... within 0-100 kRad), so we walk from the narrowest curve to the widest and accept the first curve whose own dose estimate lands inside its validity range:

for curve in [0-1, 0-5, 0-10, 0-50, 0-100] kRad:   # narrow -> wide
    dose = (dVt / curve.A) ^ (1 / curve.B)
    if dose <= curve.dose_max:
        return curve, dose          # most accurate valid fit
# if none qualify, report the 0-100 kRad estimate, flagged EXTRAP

A reading whose dose exceeds even the 0-100 kRad fit is flagged as an extrapolation rather than silently trusted.

# 6. Data-quality control

The real telemetry contains corrupted rows. We never silently drop them; we flag them and exclude them from statistics so the dose reflects genuine sensor behavior. A reading is excluded when:

the voltage is non-finite or outside the plausible band (e.g. the ~3.8e14 corruption seen in real data);

the ADC is at/near full scale (saturation / rail hit), such as the synchronous spike where all channels read ~max for one cycle;

the timestamp is out of range (e.g. 1969-epoch rows). Such timestamps are treated as missing so they cannot masquerade as the first reading; the voltage may still be usable.

Counts of each rejection category are reported so data quality is visible, not hidden.

# 7. Per-sensor aggregation and trial combination

For each (sensor_group, channel), average dVt over the valid readings to get the sensor's mean dVt, with its sample standard deviation and standard error of the mean (SEM = std / sqrt(n)).

Convert the mean dVt to dose using the selected curve (Section 5).

Combine the R1 and R2 dose estimates for each channel into a single per-configuration dose (their mean). Half the spread between the two trials is reported as a trial-to-trial uncertainty.

# 8. Uncertainty quantification

We propagate a 1-sigma dose uncertainty through the inversion Dose = (dVt/A)^(1/B), combining measurement noise on dVt with the calibration uncertainties sigma(A) and sigma(B):

(sigma_Dose / Dose)^2 =
      ( sigma_dVt / (B * dVt) )^2     # measurement (SEM on dVt)
    + ( sigma_A  / (B * A)   )^2     # calibration uncertainty in A
    + ( ln(Dose) * sigma_B / B )^2   # calibration uncertainty in B

The measurement term uses the SEM of the sensor's readings. We report both this propagated 1-sigma and the independent trial-to-trial spread (R1 vs R2); broad disagreement between the two is a flag for an unmodeled systematic.

# 9. Shielding-effectiveness analysis

Using the bare sensor (channel 1) as the reference, each shielded configuration is summarized by:

Attenuation factor = dose(bare) / dose(config)  (higher means more effective shielding).

Dose reduction (%) = (1 - dose(config) / dose(bare)) x 100.

These are reported per configuration alongside the combined dose and its uncertainty.

# 10. Outputs

A text report: data-quality summary, per-sensor / per-trial dose, and the shielding-effectiveness table.

A tidy CSV: one row per shielding configuration (dose, uncertainties, attenuation factor, % reduction) for downstream plotting and reporting.

# 11. Reference implementation

The sample_analysis/ folder contains a runnable demonstration on synthetic data:

calibration.py - calibration coefficients, range selection, uncertainty, and the channel->shielding map (single source of truth; this document is generated from it).

generate_fake_data.py - builds an ISS-format dataset with known true doses and injected anomalies.

analyze_sample.py - runs the full pipeline; on real data only the input path changes.

# 12. Assumptions and open questions for review

The voltage column is the threshold-voltage shift dVt (already baseline-subtracted). Please confirm.

Channel->shielding mapping (Section 4) is as listed. Confirm before the flight run.

The Co-60 ground calibration is assumed applicable to the mixed ISS radiation environment; any spectral correction factor is out of scope here and would be applied downstream.

Valid-voltage band and saturation threshold are tuned to the lead-enclosure baseline; revisit for the higher dVt expected on orbit.